FlightView Manual
Last updated: 2022-04-07 EHL

What is FlightView?

FlightView, also known as LiveView, is a program designed at JPL to provide real-time analysis of data from imaging spectrometers arriving at a computer over CameraLink. The computer uses Nvidia’s CUDA toolkit to compute the standard deviation of the image on a per-pixel basis, which is useful both for eliminating noise sources and as a high-sensitivity visualization of the data itself. FlightView is an open-source technology available via the NASA-JPL Github site: https://github.com/nasa-jpl/LiveViewLegacy

Installing FlightView

# Part 1: Prerequisites:

## OS:
FlightView has been tested on recent versions of Ubuntu Linux and Linux Mint. Most development has taken place using Linux Mint 20.2 with kernel 5.4.0-74-generic.

## GPU:
The computer must be using the proprietary Nvidia graphics drivers as well as a matching cuda driver. Additionally, a CUDA development kit must be installed which includes the nvcc nvidia cuda compiler and associated libraries. All these can be installed by using a single download.

Nvidia CUDA linux how-to:
https://docs.nvidia.com/cuda/cuda-installation-guide-linux/index.html

Be sure to remove any prior version(s) of CUDA and/or the Nvidia drivers prior to installing.

Install your kernel headers and build-essential files. You may need to install the kernel headers again if you update your OS:

sudo apt-get install build-essential
sudo apt-get install libc-dev libc6-dev gcc g++
sudo apt-get install linux-headers-$(uname -r)

It is strongly recommended to install the ssh server so that if the graphics do not come up, an ssh connection can be made for diagnostics:

sudo apt-get install openssh-server

The Nvidia CUDA download will include drivers for the nvidia card, cuda support files, and the nvcc.

It is recommended to use the “.run” shell script for the installation, and to not use the RPM or DEB options.

To run the “.run” installer file, download it to a known location that is not in /tmp. Next, log out from the GUI desktop. Press CONTROL-ALT-F2 to get to a text console, or ssh in from a local machine. Issue the following command to stop the desktop process (this is needed in order to run the installer):

sudo systemctl stop lightdm

Now add execute permission to the runfile, like this (modify for your exact filename):
chmod 755 cuda_11.6.0_510.39.01_linux.run

Now run the file, as root:
sudo ./cuda_11.6.0_510.39.01_linux.run

If you run into errors, there are two log files to examine:
/var/log/cuda-installer.log
and
/var/log/nvidia-installer.log

Be sure to keep the installer logs, and follow the directions carefully to include the library location with your system. On this particular linux workstation, the file /etc/ld.so.conf.d/cuda-11-6.conf was created with the following contents:

/usr/local/cuda-11.6/targets/x86_64-linux/lib

Do not copy this text directly; adjust it for your system. Be sure to run sudo ldconfig after updating the file.

Make sure to blacklist the open source nvidia graphics driver prior to rebooting:

sudo bash -c "echo blacklist nouveau > /etc/modprobe.d/blacklist-nvidia-nouveau.conf"
sudo update-initramfs -u
sudo update-grub

You may want to take the opportunity to blacklist the wifi or other unneeded hardware while you are editing these files.
## Camera Link drivers:
Make sure to install the latest EDTpdv drivers:

https://edt.com/file-category/pdv/
(download the .run file and follow the directions. The installer must be run as root, and the default location, /opt/EDTpdv is recommended)

Typically, the installer will attempt to build all the EDTpdv utilities. Some of these may fail, but generally the most important ones will build without too much trouble.

Recent EDT drivers need this one fix to the permissions:
sudo chmod -R o+r /opt/EDTpdv

The drivers include a kernel module, which can be manually re-built with the following commands:
cd /opt/EDTpdv/module
sudo make
(make includes install)
sudo update-initramfs -u
sudo update-grub

Rebuilding this module after OS updates is generally required as the kernel will update without bringing in the EDTpdv drivers. Always make sure you have the correct linux-headers package installed (see previous steps for the nvidia driver).

To initialize the camera link port, the initcam program is used with a camera link config file. By default, a large set of camera link config files are provided in /opt/EDTpdv/camera_config. These plaintext files may be used as a starting point for building a custom config file.

Here is an example for a 640x481 14-bit camera:

camera_class: "Imaging Spectrometer"
camera_model: "Camera Link 14-bit"
camera_info: "640x481 HSI"
width: 640
height: 481
depth: 14
extdepth: 14
CL_CFG_NORM: 02
continuous:1
MODE_CNTL_NORM: 00

Here is an example for a 1280x328 16-bit camera:

camera_class: "Imaging Spectrometer"
camera_model: "FPA"
camera_info: "1280x480 16-bit HSI"
width: 1280
height: 328
depth: 16
extdepth: 16
rbtfile: aiagcl.bit
continuous: 1
CL_CFG_NORM: 00
htaps: 4
serial_baud: 19200

A Camera Link port can be initialized like this:
/opt/EDTpdv/initcam -f /opt/EDTpdv/camera_config/spectrometer.cfg

It is recommended to build a shell script for this task and then to call it from a desktop launcher, with the terminal opening to show any error messages.

Example desktop file: init_camera.desktop:

[Desktop Entry]
Name=initcam NGIS
Exec=/home/eliggett/bin/init_ngis.sh
Comment=640x481
Terminal=true
Icon=cinnamon-panel-launcher
Type=Application

Example shell script file: init_ngis.sh:

#!/bin/bash
CFGFILE=/opt/EDTpdv/camera_config/NGIS.cfg
echo "Running $0 for config file $CFGFILE ..."
/opt/EDTpdv/initcam -f $CFGFILE
echo "Done! Closing window in 2 seconds."
sleep 2;

To verify the port was initialized, run the /opt/EDTpdv/take command, which attempts to acquire one frame from the camera link port. The command will typically return one of the following:
- EDT /dev/pdv0 open failed. pdv_open(pdv0_0): No such device or address
  - This means that either the kernel module did not load, or the card has a problem or isn’t physically installed. Check the output of lsmod and dmesg for details.
- pdv0: Invalid image size. Make sure device has been initialized.
  - (Run initcam)
- 1 images 1 timeouts 0 overruns 0 bytes
  - The port was initialized, but the camera was not running.
- 1 images 0 timeouts 0 overruns
  - The port was initialized, and the camera returned a frame without timeout. This means the camera link system is ready for use.

The return value is generally “1” upon error and “0” upon success.

# Part 2: Building and Installing FlightView

## Libraries needed to build FlightView:

It is recommended to install the following packages before building FlightView:
sudo apt-get install build-essential
sudo apt-get install libc-dev libc6-dev gcc g++
sudo apt-get install linux-headers-$(uname -r)
sudo apt-get install qt5-default qt5-qmake qt5-qmake-bin
sudo apt-get install git
sudo apt-get install libboost-all-dev
sudo apt-get install libqt5svg5 libqt5svg5-dev
sudo apt-get install qtcreator

## Download, configure, and build FlightView:

To download, use the following steps:
mkdir ~/Documents/FlightView
cd ~/Documents/FlightView
git clone https://github.com/nasa-jpl/LiveViewLegacy.git -b flightview

Now update the submodules (used for the Atlans A7 binary GPS library):
cd LiveViewLegacy
git submodule init
git submodule update

Configure the Makefile for cuda_take (the GPU backend) for your graphics card:
Visit this URL and find the correct specification for your card:
https://arnon.dk/matching-sm-architectures-arch-and-gencode-for-various-nvidia-cards/
Look for the correct “compute” and “sm” codes. If the cuda version is greater than 11.6, these codes may be removed and “arch=native” may be substituted.

cd cuda_take
(current directory is ~/Documents/FlightView/LiveViewLegacy/cuda_take)

Edit the Makefile, and change the following line to match your card:
NVCCFLAGS = -gencode arch=compute_86,code=sm_86 -lineinfo -Xcompiler -rdynamic --compiler-options '-Wall -Werror -Wno-unused-function'

Verify that the version of nvcc on the $PATH is the correct version:

which nvcc
nvcc -V
nvidia-smi

The nvidia-smi command should show the same CUDA Version as the nvcc -V command. If nvcc isn’t within the current path, be sure to correct this before proceeding.

To compile the nvidia CUDA back-end, issue the following command from within the cuda_take directory:

make -j

If you receive errors about “unknown architecture” then the Makefile will need to be adjusted to specify the correct compute and sm codes. If the error persists, it may mean that your card is newer than the driver supports and you will need to find a beta version of the latest Nvidia cuda drivers.

If you receive this error:
/bin/sh: 1: nvcc: not found
Then you may need to add the nvidia cuda bin directory to your current path like this (verify this location on your system before proceeding):
export PATH=$PATH:/usr/local/cuda/bin
(On many systems, the specific cuda folder, such as “cuda-11.4” will be symlinked to “cuda” for convenience.)

Next, create a directory for building FlightView:

cd ../..
(location is now ~/Documents/FlightView)
mkdir build
cd build
(location is now ~/Documents/FlightView/build)

Run qmake to configure the Makefile:
For a release build (for deployment):
qmake ../LiveViewLegacy/liveview.pro

For a debug build (for testing):
qmake CONFIG+=debug ../LiveViewLegacy/liveview.pro

Build FlightView:
make -j

When FlightView builds, it will be located in a folder within the build directory called “lv_release”, which also contains an icon and an example plaintext desktop launcher (which will likely need to be edited slightly).

It is recommended to install the files as follows:
liveview: /usr/local/bin/liveview
liveview_icon.png: /usr/share/pixmaps/liveview_icon.png
liveview.desktop: /usr/share/applications/liveview.desktop

# Part 3: FlightView Configuration and Startup

FlightView has several command-line startup options as well as a plaintext settings file containing interface preferences.

## Command-line arguments:

The command-line options may be displayed by launching FlightView with the “-h” or “--help” arguments. Some arguments listed out may not be supported. Arguments are not case-sensitive, although specified sub-argument values, such as file paths, are case sensitive.

Useful arguments:
Argument name	Description	Example
--datastoragelocation	Specifies the location for saved raw image data and log files	--datastoragelocation /mnt/flightdisk/data
--flight	Flight mode. In this mode, the operator is not prompted for file names of recordings, and a GPS connection is required.	--flight
--no-gps	Disables GPS functionality	--no-gps
--gpsIP	Specifies the IP address of the Atlans A7 TCP/IP Binary V5 GPS data server.	--gpsIP 192.168.2.6
--debug	Enable experimental and/or debug options. Not recommended for deployment.	--debug
--no-stddev	Disable the standard deviation calculation, which uses the GPU	--no-stddev

These arguments can (and should) be included in desktop launcher icons. For example, two launchers can be created, one for flight mode and one for ground testing.

Recommended Flight command-line settings:
liveview --flight --datastoragelocation /mnt/data/flightdata --gpsIP 10.0.0.6 --no-stddev

Recommended Ground command-line settings (none):
liveview

## Plaintext Settings File:

The plaintext settings file is read from this location:
~/.config/FlightView/FlightView.conf

It is recommended to first launch FlightView, press “Preferences”, and then press “Save Settings”. This will create a basic config file which could be edited. Most of the options within the settings file can be edited from the GUI of FlightView, so it is not necessarily required to edit this file. The following is an example configuration file, with comments (do not replicate with the comments, this has not been tested):

[Camera]
brightSwap14=false ; causes 0 to be bright and 2^14 to be dark
brightSwap16=false ; same as above but for 16-bit data
nativeScale=true ; Uses the full dynamic range instead of bit masking.
skipFirstRow=true ; skips the first row for standard deviation calculations (skips metadata)
skipLastRow=false ; skips the last row for standard deviation calculations
use2sComp=true ; Converts 2s compliment data into 16-bit unsigned int.

[Flight]
hideFFT=true ; hides the FFT tab
hideHistogramView=false ; hides the standard deviation histogram tab
hideHorizontalCrosshairProfile=false ; hides the horizontal crosshair profile tab
hideHorizontalMeanProfile=false ; hides the horizontal mean profile tab
hidePlayback=true ; hides the playback tab (strongly recommended to hide)
hideStddeviation=false ; hides the standard deviation tab
hideVertCrosshairProfile=false ; hides the vertical crosshair profile tab
hideVertMeanProfile=false ; hides the vertical mean profile tab
hideVerticalOverlay=true ; hides the vertical overlay tab
hideWaterfallTab=true ; hides the monochromatic data waterfall tab (not the same as the flight waterfall)
percentDiskStop=99 ; sets the absolute maximum disk utilization level. Will not record past the indicated percent (99% in this case)
percentDiskWarning=85 ; sets the disk warning utilization threshold. 85% in this case.

[Interface]
dsfCeiling=1678 ; dark subtraction ceiling slider value
dsfFloor=-580 ; dark subtraction floor slider value
fftCeiling=1000 ; FFT ceiling value
fftFloor=0 ; FFT floor value
frameColorScheme=0 ; currently selected frame color scheme (“JET”)
frameViewCeiling=6193 ; frame view ceiling value
frameViewFloor=1678 ; frame view floor value
stddevCeiling=799 ; standard deviation ceiling value
stddevFloor=242 ; standard deviation floor value
useDarkTheme=true ; use dark theme for plots
preferredHeight=1200 ; preferred main window height in pixels
preferredWidth=1600 ; preferred main window width in pixels

[RGB]
; Flight tab waterfall color selection arrays. Data are often out of order but it seems to work fine.
bandsRGB\1\bandBlue=300 ; band (horizontal line of a frame) for blue, preset 0
bandsRGB\1\bandGreen=200 ; band (horizontal line of a frame) for green, preset 0
bandsRGB\1\bandName=Best ; Name of preset 0.
bandsRGB\1\bandRed=100 ; band (horizontal line of a frame) for red, preset 0

(continues on next page)

# Part 4: FlightView Flight Interface

The flight interface is designed to give the operator a view of the image data, GPS data, disk status, and camera link status. File naming for recordings is automatic.

## Interface items:

Tab bar: Select different tabs to see the image data differently. Tabs may be locked out using the settings file.

Left-side image: Waterfall image. The selected red, green, and blue bands are used to determine pixel RGB values (see the RGB sliders). The waterfall may be expanded left-to-right by dragging the divider at the right edge.

Frame view: This area (cyan in this particular image) shows the raw frame being recorded, optionally with the data being first subtracted from a dark reference. The color scheme (a LUT between brightness values and colors) may be selected in the Preferences window. The scroll wheel can be used to zoom in and out of the image. Panning (once zoomed in) can be done by dragging the image around. Zoom may be restricted to X or Y only with the checkboxes. Double-click to set a “crosshair” from which the crosshair profiles and monochromatic waterfall (hidden) may be drawn from. The crosshair display can be disabled with the checkbox “Display Crosshairs on Frame”.

Line Plot: Shows the Roll and Pitch GPS data. The plot is intended as a “sanity check” so that operators know if the data seem sensible. The divider between the frame view and the line plot may be dragged up and down to change the division of the available space.

Flight data information area: The space below the line plot contains several status indicators as well as basic GPS data.

Clear Errors: This button clears stored errors. Press this button when an error has been acknowledged by the operator and resolved. (See the “Log” for more error information.) Program operation is not altered by this button; it merely clears the error indicators. If an error is persisting, the button may appear not to work, so be sure to check the log for details.

GPS Status: This “LED” has three states:
- Red X: Error Status. The error may be any of the following:
  - GPS dropped connection
  - Bad decode (such as a bad checksum). One bad decode at initial start is common.
  - Lack of GPS Time telegram (expected once per second)
- Blue circle: Initial status: The GPS is not connected when blue, and has not been rejected or dropped. This is also the status when the --no-gps argument is passed in to flightview
- Green circle: This indicates the GPS connection is healthy, time messages are being received, and the messages are decoding correctly.

CameraLink Status: This “LED” has two states:
- Red X: Frame rate has fallen below 2 FPS. Indicates the camera is likely not producing data.
- Green circle: Frame rate exceeds 2 FPS.
Generally, the EDT driver will time out and produce about 1.5 FPS if there is not any data coming in to the camera link ports.

Disk: This “LED” has three states:
- Green: Disk utilization is below the warning threshold as set by the settings file
- Yellow: Disk utilization is at or above the warning threshold, but below the error threshold
- Red: Disk utilization is at or above the error threshold set by the settings file. The data are not
  recorded once this happens.
See the Log and also the Disk “progress bar” meter

GPS Latitude and Longitude: Decimal degrees of latitude and longitude. “########” is the initial state prior to receiving data.

GPS Altitude: GPS Altitude, referenced to MSL, in feet.
Heading: GPS Heading in degrees. Readings are from 0 to 360 degrees with decimals.

UTC Date: System date from the GPS unit
UTC Time: Time message received once per second from the GPS unit
UTC Validity: Validity time received with every GPS message (200 times per second). As much as one second difference between the UTC Time and UTC Validity is acceptable and common.

Record Dark Frames: Begins recording a reference frame into a memory buffer. The frames are averaged together in a 64-bit floating point reference frame.

Stop Dark Frames: Press this button to end recording dark frames. Typically 100 to 1000 dark frames are recorded for a reference.

FPS @ backend: This is the framerate received into the cuda back end.

Load Dark Mask: Allows the operator to load a pre-recorded dark mask.

Server IP: This is the ip address that the internal flightview server is using. The TCP/IP server operates on the listed Server Port and accepts connections with the following commands:
- Save to file
- Stop saving
- Query FPS
The intended use of the Server is to automate long data collection sequences. Sample client code, both for python and c++, is provided in the nasa-jpl github under “utils”

Precision slider: Limits the range of the floor and ceiling sliders to -2000 to 2000, useful for when the data are dark subtracted.

Apply Dark Subtraction Filter: Causes the displayed data to be first subtracted from the dark reference data.

Ceiling and Floor sliders: The RGB conversion is clamped to the ceiling and floor selected here and scaled accordingly. Use these sliders to “zoom in” on the most useful range of the instrument’s data.

Red, Green, and Blue band sliders: These sliders select the band (horizontal column) of frame data to be used for the RGB waterfall display.

WF Length: Adjusts the apparent “length” of the waterfall display.

Preset Drop-down combo box: Use this drop-down to select a preset number. The presets contain red, green, and blue band values and a name. Select “Rename…” to give the preset a meaningful name, for example “Calibration lamp” or “Water Vapor”.

Save Preset: Press this button to save the presets to the settings file.

Save Frames: Press this button to start saving frames. In Flight mode, the frames are named automatically as follows:

Raw data: YYYY-MM-DD_hhmmss-scenedata.raw
Data header: YYYY-MM-DD_hhmmss-scenedata.hdr
Binary GPS data: YYYY-MM-DD_hhmmss-scenegps.bin

Additionally, the following two files are created once per FlightView session:
FlightView event log: YYYY-MM-DD_hhmmss-FlightView.log (same as Log dialog box)
Primary GPS data: YYYY-MM-DD_hhmmss-gpsPrimary.bin (contains a copy of all GPS data received during the session)

Note: GPS data may be “decoded” using the gpsGUI utility and libraries, available here:
https://github.com/nasa-jpl/gpsGUI
This utility can also connect to an Atlans A7 and display the data in real-time. Important: The Atlans A7 only supports one TCP/IP client connection at a time, so do not run the gpsGUI at the same time as FlightView).

Save Location: Only in ground (non-flight) mode. This button brings up a standard save dialog box where the user can select a location for saving files.

Stop Saving: Stops saving raw data.

Total #Frames / # Averaged: These two boxes contain the total number of frames recorded for a given raw data collection, and, if used, the number of frames averaged together. In non-flight mode, the operator may enable averaging, for example, averaging every 10 frames together. This is useful in long data collections where the scene is going to be averaged later anyway, such as ground calibration data. The format for average data collections is 32-bit float.

Data Location: The location of the saved data as specified with the --datastoragelocation argument. In non-flight mode, the operator may edit this text.

Disk progress bar: Shows the percent disk utilization. Placing the mouse on top of the progress bar shows the amount of space available in GB. Note: System startup disks contain an additional 5% margin which cannot be recorded to. The progress bar takes this into account automatically, thus, 100% full may actually be only 95% full on a startup disk.

Debug: This button should not be pressed and is only visible in debug builds. It is used to aid developers in testing functions without having to build UI elements for every test. Debug builds should not be used in deployments and this button should not be pressed.

Preferences: Opens the preferences dialog box, from which the settings may be changed and saved.

Camera Type: Indicates the type of camera detected. This is depreciated and will not be maintained.

Frame resolution: Frame geometry, not including any metadata rows

Enable 2s compliment filter: Converts 2s compliment data into 16-bit unsigned integer.

Display data from: Sets the limits on data which would be applied if the 14-bit bright-dark swap is used.

Native Scale: Specifies that the data shall be left alone.

Bright Swap: These buttons cause the data to be mapped to an inverted scale. The 16 bit swap is direct; the 14-bit swap uses the above “display data from” values.

Image Color Scheme: Selects from the built-in color lookup tables.

Use dark theme: Switches the frame and line plots to a dark theme. To make the overall UI of FlightView match the dark theme, select dark theme elements in the Desktop Environment’s theme selector control panel.

Save Settings: Causes the settings, including other UI-related settings such as the RGB presets, to be saved to the settings file. The settings are not automatically saved, the user must press this button to save the settings.

Close: Closes this preference dialog box

Log: Opens the Console Log dialog box, where program events, including errors, are listed. The log is saved to a unique (time-based) name in the location specified in --datastoragelocation. The filename has the following format: YYYY-MM-DD_hhmmss-FlightView.log. The log dialog box also allows the operator to append notes for later usage, for example, interesting data or malfunctions. The file is ascii text.